## Slide 1: Spatial Distortion in EPI Sequences (BOLD or DWI)
Sampling rate of EPI sequence (e.g., BOLD or DWI) in phase encoding direction (PED) is slow, which makes spatial distortion artifacts in EPI image in PED
Spatial distortion could be corrected by using B0/spatial field mapping, e.g., “gre field mapping” or “spin-echo field mapping”
  Field mapping MUST be acquired under same B0 shimming condition
  FoV and Adjustment volume MUST be identical those as BOLD/DWI acquisition
  Field maps acquired with different B0 shim condition from BOLD/DWI CANNOT be applied for distortion correction
Typically, BOLD/DWI and Field map sequences are grouped, for example,
  fMRI, SpinEchoFieldMap_AP, and SpinEchoFieldMap_PA
  fMRI, and GREFieldMap
  fMRI, SpinEchoFieldMap_AP, and SpinEchoFieldMap_PA, and GREFieldMap
  DWI_AP, and DWI_b0_PA
    Note DWI_AP includes SpinEchoFieldMap_AP, and DWI_b0_PA is SpinEchoFieldMap_PA

## Slide 2: Example 1: BOLD fMRI
BOLD fMRI data; rsfMRI, fMRI1, fMRI2
  Same FoV and Adjustment volume assigned for all three BOLD sequences
Spin Echo Field Map acquired for distortion correction
  FoV and Adjustment Volume MUST be same as those of BOLD
  B0 shim values MUST be same as those of BOLD
  Pair of forward and backward phase encoding direction (PED)
, SpinEchoFieldMap (forward & backward)
Same shim values applied -> Correct scan!

## Slide 3: Example 2: Diffusion Weighted Imaging (DWI)
DWI acquisition
  Main DWI data acquired w/ multiple b value and diffusion gradient table
  PED P > A
  DWI1_PA, DWI2_PA
b0 reference
  Same FoV and Adjustment volume as DWI data
  Same shim values as in DWI acquisition
  PED A > P
  DWI1_b0_AP, DWI2_b0_AP
Same shim values applied -> Correct scan!

## Slide 4: Example 3: BOLD fMRI
fMRI data
  BOLD acquisition given FoV and Adjustment volume
Field mapping
  Spin echo field mapping with forward-backward PED
  GRE field mapping
  Same FoV and Adjustment volume as in BOLD fMRI scan
Success!
Failure!